## Slide 1: How to Update Firmware on Respimatic 100
Step-by-step Procedure

## Slide 2: Equipment needed
Windows Laptop
USB Cable
Respimatic 100
One end with
micro-USB connector

## Slide 3: Firmware Update Procedure Overview
One-time Download Arduino Builder
One-time Install Arduino Builder
Install Respimatic Firmware Release
Download Respimatic Firmware Release

## Slide 4: Download Arduino Builder

## Slide 5: Download Arduino Builder Step 1
Open URL https://www.respimatic.com in your browser
Click on
“Update System Firmware”

## Slide 6
Download Arduino Builder Step 2
Click on
“One-time Download
Arduino Builder”

## Slide 7
Download Arduino Builder Step 3
The .exe is downloaded to your Downloads folder
Read and Dismiss the information popup

## Slide 8: Install Arduino Builder

## Slide 9: Install Arduino Builder Step 1
Find the downloaded .exe file (in your Downloads folder)
Double Click on the .exe file to execute it
You will need admin privileges on your laptop to execute this file
Your laptop’s anti-virus software will sanitize the file automatically

## Slide 10: Install Arduino Builder Step 2
The Arduino Builder installer will guide you step-by-step
Click on “Next”

## Slide 11: Install Arduino Builder Step 3
Click on “Install”

## Slide 12: Install Arduino Builder Step 4
Click on “Finish”

## Slide 13: Install Arduino Builder – Step 4
Arduino Builder is installed!
You can now delete the .exe file in your Downloads folder

## Slide 14: Download a Firmware Release

## Slide 15: Download Firmware Release Step 1
Open URL https://www.respimatic.com in your browser
Click on
“Update System Firmware”

## Slide 16
Download Firmware Release Step 2
Click on the download icon button next to the release you wish to download

## Slide 17
Download Firmware Release Step 3
A .zip is downloaded to your Downloads folder
Read and Dismiss the information popup

## Slide 18
Download Firmware Release Step 4
Double click on the downloaded .zip file to open it
Click on the “Extract All” button

## Slide 19
Download Firmware Release Step 5
A folder with the same name as the release tag will be created
It should show two files within it
Dashboard.ino.nodemcu.bin
Respimatic100.ino.mega.hex
Release is now downloaded and ready to be uploaded to the Respimatic system

## Slide 20: Install Firmware Release

## Slide 21: Install Firmware Release Step 1
Search for “Arduino Builder” on the laptop

## Slide 22: Install Firmware Release Step 2
Click on
“Arduino Builder”
in the search window

## Slide 23: Install Firmware Release Step 3
Arduino Builder app window will appear

## Slide 24: Install Firmware Release Step 4
Connect laptop’s USB port to the micro-USB port labelled “Controller” on the Respimatic system

## Slide 25: Install Firmware Release Step 6
Select “Arduino Mega 2560/ADK” from the dropdown menu in the “Target” field
Check “Verify” box

## Slide 26: Install Firmware Release Step 7
Click “Load Sketch/Binary” button on the Arduino Builder
Select Respimatic.ino.mega.hex from the extracted folder from the downloaded release

## Slide 27: Install Firmware Release Step 8
Check “Refresh” button on the
Arduino Builder
It refreshes the list of COM ports that the Respimatic system could be connected to

## Slide 28: Install Firmware Release Step 9
Click the correct COM port on the Arduino Builder
i.e. the laptop port that the Respimatic system is connected to (in this example it is COM3)
Arduino Builder will now install the selected file on the Respimatic system
Now we need to follow a similar procedure for the other file in the release.

## Slide 29: Install Firmware Release Step 10
Connect laptop’s USB port to the micro-USB port labelled “Wi-Fi” on the Respimatic system

## Slide 30: Install Firmware Release Step 11
Select “NodeMCU (ESP8266)” from the dropdown menu in the “Target” field
Check “Verify” box

## Slide 31: Install Firmware Release Step 12
Click “Load Sketch/Binary” button on the Arduino Builder
Select Dashboard.ino.nodemcu.bin from the extracted folder from the downloaded release

## Slide 32: Install Firmware Release Step 13
Click the correct COM port on the Arduino Builder
i.e. the laptop port that the Respimatic system is connected to (in this example it is COM3)
Arduino Builder will now install the selected file on the Respimatic system
Now we need to follow a similar procedure for the other file in the release.

## Slide 33: DONE